Date: xx March 2012

FIELD SAFETY NOTICE

Commercial name of the product:	Abdominal Swabs, X-Ray Detectable 30x40cm

Type of action:				Return of the Device

Attention:				Theatre Manager, Distributor

Details on affected devices:

| Batch Number | Product Code |
| --- | --- |
| 12029812 | 185460-08 |
| 12064159 | 185460-08 |
| 12039823 | 185460-08 |

Description of the problem:

It has been identified that the material of the Abdominal Swabs, X-Ray Detectable, in some instances, may tear in part or completely along the area of placement of the XRD material which is incorporated into the gauze swab. As such, we are advising recall of products with the affected batch numbers of this device as identified above, as a precautionary measure.

Actions to be taken by the user:

- Please identify and isolate all affected unused Abdominal Swabs, X-Ray Detectable products at your facility.

- Please complete the attached Confirmation form/fax back per its instructions. This step is required to confirm receipt of communications with all customers.

- Return the attached response form even if no recalled product is in inventory.

- If you have forwarded any affected product to any other healthcare institutions, please forward a copy of this letter and fax back containing affected serial numbers to those institutions.

- Mölnlycke Health Care will arrange for collection of the product from your facility.

Please contact your local Molnlycke Health Care Customer Service or Account Manager if you have any questions or concerns regarding this notification. You may also contact:

Vigilance: Caroline Price (vigilance@molnlycke.com) or. +44 (0)161 777 2646

Molnlycke Health Care also confirms that this notice has been notified to the appropriate Regulatory Agency.  Please be assured that maintaining a high level of safety and quality is our highest priority. If you have any questions, please contact us immediately.

Sincerely,

Caroline Price		   Steven Dowdley

Vigilance Associate	   Global Director of Regulatory Affairs

PLEASE COMPLETE AND RETURN THIS FORM TO:

Caroline Price, Vigilance Associate

Mölnlycke Health Care

2 Omega Drive

Irlam

Manchester

M44 5BJ

Tel: +44 (0)161 777 2646

Fax: +44 (0)161 621 2045

E-mail: vigilance@molnlycke.com

Ref - 50037042

| Product code | Batch/LOT | Quantity Quarantined (pieces/trays) |
| --- | --- | --- |

NAME : ____________________________________________________________________

POSITION : _________________________________________________________________

EMAIL ADDRESS : ___________________________________________________________

SIGNATURE : _______________________________________________________________

DATE : _____________________________________________________________________

HOSPITAL/INSTITUTE : _______________________________________________________

HOSPITAL CONTACT TELEPHONE NUMBER : ____________________________________

COUNTRY : _________________________________________________________________