Urgent Field Safety Notice

AK 98

FA Number: FAV-2024-006

Manufacturer: Gambro Lundia AB (SRN SE-MF-000018119)

Type of Action: Correction

Month DD, YYYY (to be adapted locally)

Dear Dialysis Provider: to be adapted locally)

Baxter Healthcare Corporation (Baxter) wants to inform you that it is in the process of implementing a Correction for the AK 98 hemodialysis machines listed below. These devices are manufactured using peroxide-cured silicone tubing within the hydraulic circuit.

The Baxter AK 98 dialysis machine is intended to be used for intermittent hemodialysis and/or isolated ultrafiltration treatments of patients with chronic or acute renal failure or fluid overload.

Baxter is aware of several recalls by other manufacturers related to the potential risk of exposure to non-dioxin-like (NDL) polychlorinated biphenyl acids (PCBAs) and NDL polychlorinated biphenyls (PCBs) when using certain peritoneal dialysis and hemodialysis devices. The source of the NDL PCBAs and/or NDL PCBs in those recalls was due to the manufacturing process of the silicone tubing, which used a chlorinated peroxide initiator.

Baxter is in the process of evaluating whether these same risks are present with the AK 98 machines. Although our internal testing is ongoing, at this time Baxter does not have data to definitively conclude whether there is a safety risk. Therefore, Baxter is informing you of the potential patient safety risk while our evaluation is in process.

While this evaluation is ongoing, Baxter is also in the process of transitioning certain components in the AK 98 machines from peroxide-cured silicone tubing to platinum-cured silicone tubing. Available information indicates that NDL PCBAs and NDL PCBs are not detected in medical devices with this modified version of silicone tubing. The exact timing of this transition will vary by geographic region, but Baxter will continue to make the existing configuration of AK 98 machines available in your country until this transition occurs. Once the transition timing is available, Baxter will communicate further.

Affected Product (to be adapted)

| Product Code | Product Description | Serial Numbers |
| --- | --- | --- |
| 115244 | AK 98, 230V, EFFICIENT | All serials with runtime less than 500 hours |
| 115248 | AK 98, 230V, BIO | All serials with runtime less than 500 hours |
| 115249 | AK 98, 115V, BIO | All serials with runtime less than 500 hours |
| 115250 | AK 98, 230V, SELF-CARE | All serials with runtime less than 500 hours |
| 955403 | AK 98 V2 230V | All serials with runtime less than 500 hours |
| 955404 | AK 98 V2 230V, SELF-CARE | All serials with runtime less than 500 hours |
| 955406 | AK 98 V2 115V, BIO VERSION | All serials with runtime less than 500 hours |
| 955603 | AK 98V3 230V BIO VERSION | All serials with runtime less than 500 hours |
| 955604 | AK 98V3 230V SELF-CARE | All serials with runtime less than 500 hours |
| 955989 | AK 98 AK98 V3RU | All serials with runtime less than 500 hours |

Hazard Involved

Polychlorinated biphenyls are persistent organic pollutants that have a negative impact on the ecosystem and all living beings and continue to represent a serious risk to human health. Baxter is in the process of evaluating whether these risks are present with the AK 98 machines. The risks include neuropsychological, neurobehavioral deficits, dementia, immune system dysfunctions, cardiovascular diseases, cancer and harmful effects on the reproductive system.

Available information indicates that NDL PCBA and NDL PCB levels will decrease over time with machine use.

To date, Baxter has not received any complaints related to this issue.

Actions to be Taken

Baxter is actively working on the transition to the newly manufactured platinum-cured silicone tubing components in the AK 98 machines, and in the meantime, we are recommending the following actions:

- Do not stop dialysis treatment to patients who need it. When the new platinum-cured silicone tubing components for AK 98 machines become available, Baxter will contact you to determine whether there is a need to correct or replace your existing devices. In the meantime, we recommend you continue to use your available devices while noting that machines with greater runtime hours will have less NDL PCBA and NDL PCB levels; therefore, if you have multiple AK 98 devices available, you may consider using machines with higher runtime hours before using newer machines. NOTE: In the United States, customers are being advised to consider using machines with greater than 500 hours of runtime where they have that availability. Instructions on how to find the total runtime hours are included in Attachment A.

- Complete the enclosed customer reply form and return it to Baxter by either scanning and e-mailing it to (insert local contact information) or sending it by post to (insert local contact information), even if you don’t have any inventory. Returning the customer reply form promptly will confirm your receipt of this notification and prevent you from receiving repeat notices.

- If you purchased this product from a distributor, please note that responding via the Baxter customer portal is not applicable. If a response is requested by your distributor or wholesaler, please respond to them  according to their instructions.

- If you distributed this product to other facilities or departments within your institution, please forward a copy of this communication to them.

- If you are a dealer, wholesaler, distributor/reseller, or original equipment manufacturer (OEM) that distributed any affected product to other facilities, please notify your customers of this notification in accordance with your customary procedures. (to be adapted locally)

Further Information and Support

For general questions regarding this communication or any product issue you are experiencing, contact Baxter at (insert local contact information), between the hours of (insert local information).

The local Ministry of Health (MOH) has been notified of this action. (to be adapted locally)

We apologize for any inconvenience this may cause you and your staff.

Sincerely,

Name (to be adapted locally)

Title (to be adapted locally)

Baxter Healthcare Corporation (to be adapted locally)

Enclosures:	Baxter Customer Reply Form

Attachment A: How to Identify Runtime Hours on AK 98

Peritoneal Dialysis Patient Letter (to be adapted locally)