FIELD SAFETY NOTICE

Agilia Connect VP Range – False Air-in-line Alarm

Date: July 18, 2025

Subject: Field Safety Notice regarding false air alarms occurring on Agilia Connect VP pumps with software versions 4.2 or 4.3

Affected Product Code and Serial Numbers:

| Product | Product Code | Affected Serial Numbers |
| --- | --- | --- |
| Agilia Connect VP | Z0195XX | All |
| Agilia Connect VP MC | Z0196XX | All |
| Agilia Connect VP MC WIFI | Z0197XX | All |

Note: Agilia VP MC WIFI US (Z019735) pumps are excluded from this Field Safety Notice since they cannot use software versions 4.2 or 4.3.

Dear Customer / Health Professional,

Air-in-line alarms are an essential safety feature for users of volumetric pumps in order to prevent air from entering the patient's line. Since the release of Agilia Connect pump software versions 4.2 and 4.3, Fresenius Kabi has received reports of false air-in-line alarms.

Our investigations revealed that the new software versions 4.2 or 4.3 are more sensitive to signals detected by the air sensor during patient movement or when external forces are applied to the installed tube set (e.g. when the Healthcare Practitioner or patient pulls on the set). This higher sensitivity may trigger false air-in-line alarms. False air alarms could potentially lead to therapy delays if they occur before the infusion, or to underdosing of medication if they occur during the infusion.

Fresenius Kabi has not received any reports of patient harm related to this issue.

To reduce the risk of false air-in-line alarms in Agilia pumps operating on software versions 4.2 or 4.3, Fresenius Kabi requests:

- When relocating pump stations, always move the rolling stand itself, rather than pulling on the tube sets.

- When installing tube sets, use the hook on the left side of the pump housing to secure the tube set and keep it in place throughout the infusion.

If an air alarm occurs, follow the standard procedure to verify the presence of air in the infusion line. If no air is detected, remove the tube set from the pump, install it again, and then restart the infusion normally.

Fresenius Kabi is currently developing a new software patch for versions 4.2 and 4.3 to remove these potential false air-in-line alarms.

PLEASE COMPLETE THE ENCLOSED “URGENT FSCA RESPONSE FORM” AND SEND IT BACK TO US IMMEDIATELY AT:

E-mail: info-dk@fresenius-kabi.com

Kindly assure within your organization that all relevant persons are informed about this letter and the actions as described.

Fresenius Kabi is committed to delivering the highest standards of service, product quality,

and reliability. We sincerely thank you for your understanding and cooperation, and we apologize for any inconvenience.

If you have any further questions concerning the FSCA please contact: Mattias Afzelius, Business Unit Director Medtech.

Sincerely,

Sumayah Awees

RA Officer, Complaint Officer og dNSO

Fresenius Kabi Denmark

FIELD SAFETY NOTICE FORM

Agilia Connect VP Range – False Air Alarms

SECTION A: Hospital / Facility Details

Please fill out the information below and send the completed form to Fresenius Kabi.

| Name of Hospital / Facility: |  |  |
| --- | --- | --- |
| Hospital / Facility Address: |  |  |
| Telephone Number: |  |  |

SECTION B

I have read and understand the information provided in the letter.

| Signature: |  |
| --- | --- |
| Date: |  |